## Healthy Overview
RECORDING PARAMETERS
Parameter | Value
Subjects | 26 healthy (SUB_01–SUB_26), data collection ongoing
Sessions per subject | 1–2
Paradigm | AO+MI: Subject imagines walking while watching Walker actually walk
Trials per session | 150 total = 60 imagine + 60 walk + 30 rest (30 cycles × 5)
Session duration | ~57 min continuous recording
EEG | actiCHamp, 63ch + FCz ref, 1000Hz, BrainVision format
Goniometer | Biometrics Ltd, 1000Hz, 23ch + 1 Stim = 24ch, UTF-16LE .txt
Goni receivers | 3 DataLITE receivers × 8ch each; two configs (Type A/B) swap one channel
EEG-Goni sync | IOI cross-correlation of marker onset times (no periodic S10 triggers)
Alignment residual | < 50ms across all 13 sessions verified
EEG MARKERS — STANDARD (12/13 sessions)
EEG Marker | Meaning | Count | PyLog equivalent | Notes
S 11 | Experiment START | 1 | S11
S 12 | Experiment END | 1 | S12
S 1 | Imagine trial START | 60 | S1 (backward) + S6 (forward) | Odd/even direction collapsed into same marker
S 2 | Imagine trial STOP | 60 | S2 + S7
S 4 | Walk trial START | 60 | S4 (forward) + S8 (backward) | Odd/even direction collapsed into same marker
S 5 | Walk trial STOP | 60 | S5 + S9
S 7 | Rest START | 30 | S20
S 8 | Rest STOP | 30 | S21
— | TOTAL markers/session | 302 |  | Must match goni Stim onset count for alignment
GONIOMETER — TWO RECEIVER CONFIGS
Config | Sessions (7+6=13) | Missing channel | Present instead | Subject X (6 joints) | Walker X
Type A | SUB_01_s01, 02_s01, 02_s02, 03_s01, 04, 05, 06_s01 | LAnkW X | LAnkS Y | All 6 ✓ | 5 only (no LAnk)
Type B | SUB_01_s02, 03_s02, 06_s02, 07, 08, 09 | LAnkS Y | LAnkW X | All 6 ✓ | All 6 ✓
⚠ Channel column ORDER varies per subject (e.g. SUB_09 completely different). NEVER hardcode indices → use find_goni_idx.m

## Healthy Sessions
ALL HEALTHY SESSIONS — from Goni QC v3 + recording logs
Subject | Session | Goni file(s) | Total trials | MI | Walk | Rest | Goni problems | Status / Notes
SUB_01 | sess01 | 1 file | 150 | 60 | 60 | 30 | 0 | Clean. Old experiment script: no S7/S8, rest = R1+S4→R1+S5, 2× S11. Trial extraction needs special handling.
SUB_01 | sess02 | 1 file | 150 | 60 | 60 | 30 | 0 | Clean.
SUB_02 | sess01 | 1 file | 150 | 60 | 60 | 30 | 0 | Clean.
SUB_02 | sess02 | 1 file | 150 | 60 | 60 | 30 | 0 | Clean.
SUB_03 | sess01 | 1 file | 150 | 60 | 60 | 30 | 0 | Clean.
SUB_03 | sess02 | 1 file | 150 | 60 | 60 | 30 | 0 | Clean.
SUB_04 | sess01 | 1 file | 150 | 60 | 60 | 30 | 0 | Clean.
SUB_04 | sess02 | 2 files (split!) | 36+115=151 | 61 | 60 | 30 | 0 | Recording interrupted → restarted. Two separate goni files (36 trials then 115 trials). Must load and merge both. Trial counts: segment1=36, segment2=115.
SUB_05 | sess01 | 1 file | 150 | 60 | 60 | 30 | 0 | Clean.
SUB_06 | sess01 | 1 file | 150 | 60 | 60 | 30 | 0 | Clean.
SUB_06 | sess02 | 1 file | 150 | 60 | 60 | 30 | 0 | Clean. (v2 QC flagged trial 6, v3 reclassified as OK)
SUB_07 | sess01 | 1 file | 150 | 60 | 60 | 30 | 0 | Clean.
SUB_07 | sess02 | 1 file | 150 | 60 | 60 | 30 | 0 | Clean.
SUB_08 | sess01 | 1 file | 150 | 60 | 60 | 30 | 0 | Clean.
SUB_09 | sess01 | 1 file | 150 | 60 | 60 | 30 | 0 | Clean. NOTE: goni channel column order completely different from all other subjects.
SUB_09 | sess02 | 2 files (split!) | 15+140=155 | 6+56 | 6+56 | 3+28 | 0 | Recording interrupted → restarted. Two goni files (15 then 140 trials). Must load and merge both.
SUB_10 | sess01 | 1 file (partial?) | 20 | 8 | 8 | 4 | 0 | Only 20 trials recorded (likely equipment issue early in session). Segment 2 data unclear.
SUB_10 | sess02 | 1 file | 150 | 60 | 60 | 30 | 0 | Clean.
SUB_11 | sess01 | 1 file | 150 | 60 | 60 | 30 | 0 | Clean.
SUB_12 | sess01 | 2 files (split!) | 78+75=153 | 32+30 | 31+30 | 15+15 | 19+19=38 | SEVERE: Walker goniometer wireless disconnection → zero readings in many trials. 38 trials affected across both segments. Recording also split into 2 goni files.
SUB_12 | sess02 | 1 file | 150 | 60 | 60 | 30 | 0 | Clean.
SUB_13 | sess01 | 1 file | 150 | 60 | 60 | 30 | 9 | MODERATE: 9 trials with goni zeros (wireless disconnection). Exclude affected trials.
SUB_14 | sess01 | 1 file | 150 | 60 | 60 | 30 | 0 | v3 QC: 0 real problems (v2 had flagged 25 but reclassified as expected flat readings during rest/stand).
SUB_15 | sess01 | 1 file | 150 | 60 | 60 | 30 | 0 | Clean.
SUB_15 | sess02 | 1 file | 150 | 60 | 60 | 30 | 0 | Clean.
SUB_16 | sess01 | 2 files (split!) | 113+40=153 | 46+16 | 45+16 | 22+8 | 0 | Recording interrupted → restarted. Two goni files (113 then 40 trials). Must load and merge both.
SUB_17 | sess01 | 1 file | 150 | 60 | 60 | 30 | 0 | Clean.
SUB_18 | sess01 | 1 file | 149 | 59 | 60 | 30 | 0 | Clean. 149 trials (1 trial lost, likely marker issue).
SUB_19 | sess01 | 1 file | 150 | 60 | 60 | 30 | 0 | Clean.
SUB_20 | sess01 | — | — | — | — | — | — | Not in QC v3 output. Data status unclear.
SUB_21 | sess01 | — | — | — | — | — | — | Not in QC v3. Earlier QC flagged 19/60 GI trials with Walker goni zeros.
SUB_22 | sess01 | 3 files (split!) | 78+3+70 | — | — | — | — | SEVERE: Goni wireless disconnect → mid-session restart → 3 recording segments. Segment 2 (3 trials) has NO corresponding EEG S11 marker. Visual alignment done by Neethu. Custom loading script needed.
SUB_23–26 | — | — | — | — | — | — | — | Not yet in QC. Data collection ongoing.
SPLIT SESSION DETAIL — sessions where 1 recording became 2+ files
Subject | Session | Segment | Goni filename (timestamp) | Trials | MI | Walk | Rest | Why split / What to do
SUB_04 | sess02 | 1 of 2 | 11Mar2026-141250.957 | 36 | 15 | 14 | 7 | Recording interrupted (unknown cause). Load as separate .mat, align independently.
SUB_04 | sess02 | 2 of 2 | 11Mar2026-143125.403 | 115 | 46 | 46 | 23 | Main segment. Merge both for full session.
SUB_09 | sess02 | 1 of 2 | 23Mar2026-135433.557 | 15 | 6 | 6 | 3 | Recording interrupted. Short first segment.
SUB_09 | sess02 | 2 of 2 | 23Mar2026-142441.410 | 140 | 56 | 56 | 28 | Main segment. Merge both.
SUB_10 | sess01 | 1 of 2? | 06Mar2026-140341.688 | 20 | 8 | 8 | 4 | Only 20 trials. Possible second segment not yet documented.
SUB_12 | sess01 | 1 of 2 | 17Mar2026-105800.701 | 78 | 32 | 31 | 15 | Both segments have Walker goni zeros (19 problems each). Wireless disconnect persisted across restart.
SUB_12 | sess01 | 2 of 2 | 17Mar2026-112420.855 | 75 | 30 | 30 | 15 | Same goni issue continues.
SUB_16 | sess01 | 1 of 2 | 18Mar2026-135057.650 | 113 | 46 | 45 | 22 | Recording interrupted.
SUB_16 | sess01 | 2 of 2 | 18Mar2026-143726.758 | 40 | 16 | 16 | 8 | Shorter second segment.
SUB_22 | sess01 | 1 of 3 | 30Mar2026-153954.902 | 78 | — | — | — | Normal first segment.
SUB_22 | sess01 | 2 of 3 | 30Mar2026-161720.228 (file 1) | 3 | — | — | — | PROBLEM: No EEG S11 for this segment. No automatic alignment possible.
SUB_22 | sess01 | 3 of 3 | 30Mar2026-161720.228 (file 2) | 70 | — | — | — | Neethu created duplicate dir with visual alignment. Custom loading script: data_load_eeg_pilot_feb2026_sub22_sess1.m

## Patient Overview
PARTICIPANTS
Subject | Label | Diagnosis | Sessions | Stim OFF | Stim ON | Notes
Sub01 | RESTORE2_001 (Mr. Chong) | SCI, epidural SCS | 5 (Sess01–05) | Sess01–03 | Sess04–05 | YOB 1966, Cap 56
Sub02 | RESTORE2_002 (Ms. Stella) | SCI, epidural SCS | 3 + Resting | Sess01 | Sess02–03 | Cap 56. Resting state recorded to characterize stim ON/OFF artifacts.
RECORDING PARAMETERS
Parameter | Value
EEG | actiCHamp Plus, 63ch + Cz ref (differs from healthy FCz), 1000Hz, BrainVision format
GI paradigm | Patient seated, imagines walking while watching Walker walk. 20–55 trials/session (variable by fatigue).
Session flow | S2S(×5) → GI Block 1 → [6m Walk + GI]×2 → GI Block 2 → 2-min Walk → GI Block 3
EEG-Goni sync | S10 periodic sync pulses (83–144/session) matched to Goni Stim channel rising edges
Alignment quality | Max error 43ms across all sessions. >99% match rate.
EEG MARKERS — STANDARD (Sess03 onward)
EEG Marker | Meaning | Count/session | Notes
S 10 | Periodic sync trigger | 83–144 | For alignment ONLY. Not an experiment event. Sent to both EEG and Goni simultaneously.
S 11 | Experiment START | 1 | Absent in Sess01–02 (legacy protocol).
S 12 | Experiment END | 1 | Absent in Sess01–02; also missing in Sub02_Sess03.
S 1 / S 2 | GI START / END | 20–55 | Main analysis target.
S 3 / R 1 | 6m Walk START / END | varies | Walker walks, patient watches.
S 4 / S 5 | Sit-to-Stand START / END | 5–6 | Patient performs actual movement.
S 7 / S 8 | 2-Min Walk START / END | 0–1 | Patient walks with FLOAT support. Only 3 usable segments across all 8 sessions.
S 6, S20/21 | Relax break | varies | Alternative naming for rest periods.
LEGACY MARKERS — Sub01 Sess01–02 only (before 22 Dec 2025)
Marker | Standard meaning (Sess03+) | Legacy meaning (Sess01–02) | Notes
S 1 | GI START | Eyes Close / Imagine Start | In files 0009–0012 (Sess01), S1/S2 still appear to mean GI despite legacy protocol.
S 3 | 6m Walk START | 6m Walk START | Same meaning, different context.
S 7 / S 8 | 2-Min Walk START/END | 2-Min Walk START/END | Same meaning.
S 13 / S 14 | (not used) | 6MW Pause / Resume | Only in legacy sessions.
S 11 / S 12 | Experiment START/END | (absent) | Cannot auto-trim. Use skip_trim=true or S10 bookends.

## Patient Sessions
ALL PATIENT SESSIONS — detailed
Session | Date | Stim | EEG files | Goni files | GI trials | S2S | 2MW | Issues | Handling
Sub01 Sess01 | 30 Oct 2025 | OFF | 6 files (0001,0008–0012) | Multiple (per EEG file) | 51 (from 0009–0012) | 5 (in 0008) | 1 usable (2486–2608s) | Legacy markers (no S11/S12). 0001: only goni trigger, no EEG data → discard. 0008: S2S only, goni .txt missing → discard. 0010: USB error during recording. | pop_mergeset files 0009–0012 only. skip_trim=true. Alignment: IOI sliding match needed (192s residual with naive approach).
Sub01 Sess02 | 13 Nov 2025 | OFF | 3 files (0013–0015) | 3 goni files | ~53 (18+18+17) | 0 | 0 | Legacy markers (no S11/S12). Shoulder movement during Block 2 GI → possible motor artifact. | pop_mergeset all 3. skip_trim=true.
Sub01 Sess03 | 22 Dec 2025 | OFF | 1 file | 2 goni files (goni restarted) | ~29 | ? | 0 | Bad REF electrode (Cz impedance >20kΩ) → corrupts ALL channels. clean_rawdata flags 16–38/63 channels. Patient fatigued (arrived from overseas), only 2 GI blocks. Goni restarted mid-session → 2 files (59+44 stim vs 103 S10). | SESSION ABANDONED — do not use for analysis. If used: align each goni file separately.
Sub01 Sess04 | 20 Jan 2026 | ON | 2 files (0016, 0018) | 2 goni files | 27 (13+14 in 0016) | 5 (in 0018) | 0 | Bad REF electrode (elevated bad channels). Goni wireless disconnect → gap at 335–775s in file 1, 1 stim trigger lost. File 0018 = S2S only, no GI. First session with stim ON → eSCS EM artifact in EEG. | pop_mergeset both files. Alignment fallback (brute-force) works → 82/83 matched. Process with caution (bad REF + stim artifact).
Sub01 Sess05 | 20 Feb 2026 | ON | 1 file (0020) | 1 goni file | ~45 | ? | Partial (49s only) | Stim ON (eSCS artifact present). 2MW: S7 sent but recording ended (S12) before S8 → only 49s of walk. Patient exhausted during 6m walk, only half a rep. | Standard processing. 2MW segment too short to use.
Sub02 Sess01 | 05 Jan 2026 | OFF | 1 file | 1 goni file | 29 (14+15) | ? | 1 usable (S7+120s) | Bad REF electrode (Cz >20kΩ, elevated bad ch). FLOAT battery running out → only 2 GI blocks. Second S7: only 13s before S12 → not usable. | Process with caution (bad REF). 2MW: use S7 + 120s as end time.
Sub02 Sess02 | 29 Jan 2026 | ON | 1 file (0018) | 1 goni file | 39 (13+13+13) | 6 | 1 usable (S7+120s) | Bad REF electrode (Cz >20kΩ, elevated bad ch). Different FLOAT weight percentages (90/70/50/20%) for S2S. | Process with caution (bad REF). 2MW: use S7 + 120s as end time.
Sub02 Sess03 | 02 Mar 2026 | ON | 1 file | 1 goni file | 44 | ? | 0 | S11 present but NO S12 marker → cannot auto-trim end. 1 spurious stim pulse at goni end (145 vs 144 S10). Stronger eSCS stimulation parameters in this session. | Trim from S11 to end-of-recording. Alignment: last goni stim ignored (spurious).
Sub02 Sess04 | (later) | ON (stronger) | 1 file | 1 goni file | ? | ? | ? | SEVERE: clean_rawdata at corr=0.8 flags 33/59 channels as bad. Root cause: stronger eSCS stimulation parameters → EM artifact overwhelms scalp EEG. corr=0.7 → 11 bad; corr=0.6 → 3 bad. | Requires relaxed threshold (corr=0.6 or 0.7) or exclusion from group analysis.
ALIGNMENT QUALITY — S10 ↔ Goni Stim matching
Session | S10 count | Goni Stim count | Matched | Max timing error | Match rate | Notes
Sub01 Sess01 (6 EEG files) | 83 | 83 | 83 | 43ms | 100%
Sub01 Sess02 (3 EEG files) | 83 | 83 | 83 | 43ms | 100%
Sub01 Sess03 — Goni file 1 | 103 | 59 | 59 | — | 100% | Goni restarted mid-session → fewer stim onsets than S10
Sub01 Sess03 — Goni file 2 | 103 | 44 | 44 | — | 100% | Second goni file covers rest of session
Sub01 Sess04 — EEG file 1 | 83 | 82 | 82 | — | 99% | Goni gap 335–775s → 1 stim trigger missing in gap
Sub01 Sess04 — EEG file 2 | 83 | 83 | 83 | 41ms | 100%
Sub01 Sess05 | 83 | 83 | 83 | 41ms | 100%
Sub02 Sess01 | 144 | 144 | 144 | 41ms | 100% | Higher S10 frequency (144 vs 83)
Sub02 Sess02 | 144 | 144 | 144 | 41ms | 100%
Sub02 Sess03 | 144 | 145 | 144 | — | 99% | 1 spurious stim pulse at very end of goni recording
2-MIN WALK (S7→S8) USABILITY
Session | S7 events | S8 events | Duration | Usable? | Why / How
Sub01 Sess01 | 1 | 1 | ~122s (2486–2608s) | YES | Complete S7→S8 pair. Only session with actual S8 marker.
Sub01 Sess03 | 6 | 0 | ~20s each | NO | These S7 markers are individual short walk trials, not 2-min continuous walk.
Sub01 Sess05 | 1 | 0 | 49s (S7 to S12) | NO | Walk started but recording ended (S12) before 2 minutes elapsed.
Sub02 Sess01 (1st S7) | 1 | 0 | ~120s (2050–2170s) | YES | No S8 sent. Use S7 + 120s as endpoint (standard 2-min walk).
Sub02 Sess01 (2nd S7) | 1 | 0 | 13s (to S12) | NO | Only 13 seconds of data before recording ended.
Sub02 Sess02 | 1 | 0 | ~120s (2253–2373s) | YES | No S8 sent. Use S7 + 120s as endpoint.
RESTING STATE (Sub02 only) — separate recording, no gait markers
Segment | Start marker | End marker | Duration | Eyes | Stim | Notes
1 | eye_open_stim_off (66s) | stop (189s) | ~122s | Open | OFF | Baseline condition
2 | eye_close_stim_off (213s) | stop (330s) | ~117s | Closed | OFF
3 | eye_open_stim_on (960s) | stop (1081s) | ~121s | Open | ON | For stim artifact characterization
4 | eye_close_stim_on (1100s) | stop (1218s) | ~119s | Closed | ON